## Slide 1
FairDermNet / DermaLens
Diagnosing and Correcting Silent Fairness & Quantization Failures
in an On-Device Skin Lesion Classifier
Final-Year Engineering Project — Technical Audit & Remediation

## Slide 2
OVERVIEW
What This Talk Covers
01
System Overview
8-class taxonomy, FairDermNet architecture, on-device export pipeline
02
Three Defects Found
Stale labels, fabricated tone data, quantization collapse — all silent
03
Root-Cause Method
How each was traced, not just patched
04
Fixes & Evidence
Before/after numbers for every fix
05
What's Still Open
Retraining, real fairness gap, next steps
DermaLens — FairDermNet
2

## Slide 3
SYSTEM
What DermaLens Does
MobileNetV3-Large + Fairness Attention + GeM Pooling, dual-head, exported to TFLite for fully offline mobile inference.
8 classes
Melanoma, BCC, SCC, AK, Nevi, BKL, DF, VASC — malignant = 0–3
4 datasets
HAM10000, ISIC 2019, PAD-UFES-20 (train) + DDI (fairness held-out)
Dual-head loss
Classification + Fitzpatrick tone + equalized-odds fairness term
Target: <6 MB
INT8-quantized, on-device, no network required
Loss: L = α·L_cls + β·L_fitz + γ·L_equalized-odds
DermaLens — FairDermNet
3

## Slide 4
CONTEXT
The Trap: Good Top-Line Numbers
Training and top-1 test accuracy looked fine (73.7% balanced accuracy). Three independent defects were hiding underneath, invisible to standard accuracy monitoring.
None of the three defects below threw an exception during training.
None degraded top-1 test accuracy.
All three were found only by tracing derived values back to source, and inspecting distributions / tensor scales — not by watching accuracy curves.
DermaLens — FairDermNet
4

## Slide 5
DEFECT 1 OF 3
Stale Binary Risk-Label Mapping
WHAT WAS WRONG
class_idx in {0,1,2,3,4} left over from an old 9-class taxonomy. Class 4 = Melanocytic Nevi, the single largest class (53% of data), silently mislabeled malignant. Training loss unaffected — but DDI fairness eval, threshold calibration, and GradCAM display all consumed the corrupted label.
THE FIX
Repointed all 4 call sites to class_idx_to_binary() — the canonical taxonomy function. DDI eval switched to expert ddi_binary labels (not a derived quantity). Malignant probability corrected: sum over classes {0,1,2,3}, not a single legacy index. Result: DDI eval runs end-to-end for the first time →
AUC 0.585 · Sens 81.9% · Spec 24.1%
Equalized-odds gap = 25.96% (first reliable measurement)
DermaLens — FairDermNet
5

## Slide 6
DEFECT 2 OF 3
Fabricated Skin-Tone Supervision
Root cause of the 25% fairness gap: 97% of training tone labels were fake.
estimate_ita(img) called without its required lesion-mask argument → every call threw → exception silently caught → hardcoded fallback to Fitzpatrick III. A dataset clamp then mapped any missing tone to the same fallback.
Tone | Fabricated (before) | Estimated (after)
I | 153 | 235
II | 876 | 24,722
III (fallback) | 36,542 (97%) | 12,454
IV | 62 | 220
V | 10 | 12
VI | 1 | 1
FIX
• Pass real Otsu lesion mask to estimator • Unresolvable → explicit -1 (unknown), never faked • Loss masks -1 rows; metrics exclude -1 groups • Re-estimated all 36,150 rows • Deliberately did NOT mix in DDI (would invalidate the held-out fairness eval)
DermaLens — FairDermNet
6

## Slide 7
DEFECT 3 OF 3
INT8 Quantization Collapse
73.7% → 3.3% top-1 accuracy after full-INT8 export. Graph conversion was fine (float32 TFLite = ONNX = PyTorch, exact match).
ROOT CAUSE
GeM pooling: y = (Σ xᵢᵖ)^(1/p), learned p ≈ 2.88 The xᵖ intermediate spans ~10⁶ in magnitude → INT8 quantization scale = 13,615 (vs ~1–5 for normal conv layers). Precision is destroyed.
FIX: DYNAMIC-RANGE QUANTIZATION
INT8 weights, float32 activations — keeps GeM's power op exact. 5.51 MB — under the 6 MB on-device budget Validated on real images: ONNX 100% match · TFLite 94% top-1 / 99% clinical-decision match vs PyTorch.
DermaLens — FairDermNet
7

## Slide 8
EVIDENCE
Deployed Model vs. PyTorch Reference
150 real held-out test images, evaluated at the calibrated clinical threshold (0.25).
Chart: PyTorch vs. Deployed Model (%)
| Decision Acc. | Sensitivity | Specificity | Top-1 Acc.
PyTorch (fp32) | 73.3 | 90.9 | 63.2 | 76
Deployed (5.51 MB, dynamic-range) | 72 | 89.1 | 62.1 | 74.7
DermaLens — FairDermNet
8

## Slide 9
CALIBRATION
Clinical Decision Threshold
Swept malignant probability = Σ p(class 0–3) against the test set under a ≥96% sensitivity constraint.
0.25
Calibrated threshold
96.0%
Sensitivity
55.4%
Specificity
0.9239
AUC
Placeholder threshold was 0.5 → 77% reduction in missed malignant cases (false-negative rate 17.4% → 4.07% per malignant case) after recalibration.
DermaLens — FairDermNet
9

## Slide 10
DISCUSSION
What This Means, Beyond This Project
Retraining ≠ fixing fairness
A 25% gap rooted in corrupted labels won't shrink from more epochs. Fix supervision before spending compute.
Silent failures need targeted audits
None of the 3 defects threw errors or hurt top-1 accuracy — only tracing derived values to source caught them.
Quantization is architecture-dependent
Learned pooling exponents (GeM) can blow up INT8 activation ranges by 1000×+. Inspect per-tensor scales, not just end-to-end accuracy.
DermaLens — FairDermNet
10

## Slide 11
STATUS
Closed vs. Open Items
✅ CLOSED
• TFLite export generated
• Flutter model asset shipped (5.51 MB, on-device verified)
• Clinical threshold calibrated (0.25)
• Export parity validated (real images)
• Risk-label taxonomy bug fixed
• Tone-supervision bug fixed & re-estimated
⚠ OPEN
• Model not yet converged (~68–74% bal. acc., still rising)
• Fairness gap (25.96%) needs real dark-skin training data
• Full retrain + re-measure DDI gap
• Docs vs. implementation drift (proposal doc stale)
DermaLens — FairDermNet
11

## Slide 12
Fairness-critical ML needs
data-lineage audits, not just accuracy curves.
Thank you — questions welcome
DermaLens — FairDermNet
12

## Slide 13
Why thresholds failed: the deficit is representational
No operating point exists (per-group ROC ceilings, DDI):
  AUC by group — I–II 0.574 · III–IV 0.668 · V–VI 0.484 (CI includes 0.5) · ALL 0.585
  The malignancy score carries no measurable ranking signal for FST V–VI images.
A spurious AK response drives the false positives: the AK column is the top malignant contributor on 63.3% of benign DDI images (one-sided KS 0.76 vs in-domain benigns) and tops all 15 highest-scoring false positives.
  It is a domain effect, not a tone effect: within DDI the AK response is flat-to-highest on the lightest-tone group; at matched tone the DDI-vs-test gap is ~10×.
Post-hoc repair fails (whack-a-mole):
  per-group thresholds cannot buy specificity without paying sensitivity; AK down-weighting (λ fit on tune half only) peaks at λ*=0.25 with tune AUC 0.596→0.609 (+0.013) and collapses beyond — the argmax migrates AK→BCC/MEL.
  CAM comparators bound the mechanism: diffuseness is the AK head's normal behavior (FP 0.023 vs true-AK 0.019, MW p=0.11); field-wide placement is near-identical (edge-band 0.41 vs 0.38, p=0.28). Rulers/ink present in 8/15 FPs but attract no focal activation.
Conclusion: closing the gap requires in-domain training signal (capture-domain hard negatives + tone augmentation) — declared as future work, to be judged against the ROC ceilings.
Notes: STEP3-MECHANISM-SLIDE (auto-generated) — generated 2026-09-06T19:13:27 from model/audit artifacts (step3/3b/3c/3d reports + CSVs); same sources as tex_integrate_step3.py. Numbers: AUC V–VI 0.484 [0.382, 0.585]; AK top-contributor share on benign DDI 63.3%; KS 0.76; λ*=0.25; CAM MW p=0.11. Blinded expert review of the 15 triads pending (model/audit/blind_review_protocol.md).